Training protocol on AFM :

Atomic Force Microscopy (AFM) is a high-resolution imaging technique used to measure surface topography at the nanometer scale. It is a powerful tool for characterizing surface roughness, height profiles, and mechanical properties. The AFM operates by scanning a sharp probe over the surface of the sample, detecting forces between the probe and the sample to create a detailed 3D image of the surface.

2. System Capabilities

The AFM system offers the following capabilities:

- High-resolution surface imaging down to sub-nanometer scale.

- 3D topographical mapping of surfaces.

- Measurement of surface roughness, height variations, and defects.

- Force measurements to assess mechanical properties like stiffness and adhesion.

3. Personal Protective Equipment (PPE)

To ensure safe operation, the following PPE is required:

- Gloves: To avoid contaminating the AFM probe and samples.

4. Pre-Operation Safety Checks

Before starting the AFM measurement, perform these checks:

- Inspect the AFM Tip: Ensure the tip is clean and undamaged. A dull or broken tip will affect image quality.

- Check the Sample Stage: Ensure the sample is properly mounted and secure on the stage.

- Verify Laser Alignment: For AFMs using laser detection, confirm the laser beam is correctly aligned with the cantilever and photodetector.

- Calibration Check: Ensure the system is properly calibrated for accurate measurements, including z-axis sensitivity and scanner calibration.

5. System Operation Procedure

5.1. Start-Up Procedure

- Turn On the AFM System: Power on the AFM control system, including the electronics and the scanning unit.

- Load the Sample: Using tweezers or gloves, carefully mount the sample onto the sample stage. Ensure the sample is clean and free from dust or debris.

- Align the AFM Probe: Position the AFM tip close to the sample surface without making contact. Use the software to finely adjust the z-position of the cantilever.

- Laser Alignment: If applicable, align the laser beam onto the back of the cantilever, ensuring the reflected beam is centered on the detector.

- Engage the Probe: Use the software to lower the AFM probe until it lightly touches the surface, ready for scanning.

5.2. Scanning and Imaging Process

- Set Scan Parameters: Adjust the scan rate, scan size, and resolution to suit the sample. Start with a low scan rate to avoid damaging the tip or sample.

- Monitor Feedback Signals: Continuously monitor feedback from the system to ensure stable tip-sample interaction. Adjust parameters if needed.

- Capture and Analyze Data: Once the scan is complete, review the 3D surface image and any additional data (e.g., force curves). Analyze the surface roughness, features, or any defects present.

5.3. Shut Down Procedure

- Lift the AFM Probe: Safely retract the AFM probe using the software to prevent accidental contact with the sample surface.

- Turn Off the System: Power down the system according to the manufacturer’s procedure, ensuring the electronics and scanner are turned off.

- Remove the Sample: Carefully remove the sample from the stage using tweezers or gloves to avoid contamination or damage.

6. Safety Steps to Perform During Operation

- ESD Precautions: Always wear an antistatic wrist strap when handling the AFM tip or electronics to prevent damage from electrostatic discharge.

- Laser Safety: Do not look directly into the laser used for the AFM’s detection system. Even low-power lasers can cause eye damage.

7. Post-Operation Maintenance

- Clean the Sample Stage: Wipe down the stage with appropriate solvents (e.g., isopropyl alcohol) to remove any contamination.

- Backup Data: Save and back up all measurement data for future reference and analysis.

8. Emergency Procedures

- System Malfunction: If the system malfunctions, such as loss of laser alignment or