Safety and Work Protocol for ​​ Profilometer Dektak XT:

Only Trained and Authorized Users are Allowed to Operate the System !

- Always use only dry samples!

- Clean the sample with nitrogen prior to measurements.

- Never set the scan speed lower than 50 μm/sec; always aim for a speed around 100 μm/sec.

- Perform scans strictly within the sample limits; never start the scan outside the sample and move onto it, as this may damage the tip.

- Always position your sample on the stage while the tip is raised!

- Always remove your sample after the tip is raised!

- Reduce the light source intensity after completing your work.

- Scanning occurs in one direction, from your side to the opposite side. Keep this in mind to stay within the sample boundaries.

For any questions, call 79812/ 08-6479812 Tania

For any problems